## Test Cases
Number	Name	Description	Verification Method	Expected Results	Tolerance	Test Result	Pass	Fail	Tester ID	Test Date	Device ID	RMVS (Mandatory / Optional - M / O)
Draft v0.1 - March 25th
By Public Invention, in cooperation with BreathOfLife and Jonathan Orban
Licensed under CC0; consider a template to copy and extend
Copyright Robert L. Read, Juan E Villacres Perez, Nariman Poushin, Jonathan Orban and others 2020
Note: primary source is UK Rapid Manufactered Ventilator System (1)
Primary Point of Contact: Robert L. Read, PhD <read.robert@gmail.com>
Connection
	CON1	Electrical connection	Instruct CLI to connect system to ventilator	They can in 15 seconds	< 60 seconds		Y	N				M
	CON2	Oxygen connection	Instruct CLI to connect O2 supply to ventilator	They can in 60 seconds	< 90 seconds		Y	N				M
	CON3	Replace Batteries	Instruct CLI to replace batteries	They can in 5 minutes	< 10 minutes		Y	N				M
	CON4	Endotracheal tube connection	Instruct CLI to connect endotracheal tube (in-situ) to ventilator	They can in 45 seconds	< 60 seconds		Y	N				M
	CON5	Humidifier connection	Instruct CLI to connect humidifier to ventilator	They can in 60 seconds	< 90 seconds		Y	N				O
	CON6	Inlet and outlet separated by 10cm	Measure distance.	At least 10 cm	none		Y	N				O
	CON7	Male inlet and outlet ports provide clearance	Check inlet and outputs clear in 5cm circle from center	At least 5cm circle	none		Y	N				M
8							Y	N
9							Y	N
10							Y	N
11							Y	N
12							Y	N
13							Y	N
14							Y	N
15							Y	N
16							Y	N
17							Y	N
18							Y	N
19							Y	N
20							Y	N
Startup							Y	N
	SUP1	CLI can tell if powered	Instruct CLI to plug in, ask CLI if it is powered. Instruct CLI to plug into unconnected extension cord. Ask CLI if powered	CLI can correctly tell if powered			Y	N				M
	SUP2	Mandatory Mode on Power-Up	On initial power-up, system enters mandatory mode and begins working	Mandatory mode operating with respiration rate = DRATE and volume = DVOL and			Y	N				M
	SUP3	Volume Control begins correct	On power-up, ask CLI what volume is produced, and ask if it is being produced.	CLI answere 400 ml and V produced 400ml	10%		Y	N				M
	SUP4	Respiration Rate Standard	On power-up, check respiration rate, measure total respirations per minute	Respiration rate = DRATE, or persistent Respiration Rate, or match control setting if tied to knobs	1 rpm + 10% of setting		Y	N				M
	SUP5	I:E ratio corrects	On power-up estimate inspiration and expiration absolute times in tenths of seconds (w/ stopwatch), comptue I:E ratio	I:E ratio = DIERATIO, or remebered setting, or physical setting	none		Y	N				M
	SUP6	Mode	On power-up, in mandatory mode, or persistent mode if that is a feature	Yes	none		Y	N
	SUP7	CLI can locate instruction label	Instruct CLI to locate and inspect instruction label	CLI can correctly locate and read the instruction label in 20 seconds	30 seconds		Y	N
8							Y	N
9							Y	N
10							Y	N
11							Y	N
12							Y	N
13							Y	N
14							Y	N
15							Y	N
16							Y	N
17							Y	N
18							Y	N
19							Y	N
20							Y	N
Inspiratory Airway Pressure (Volume Control Mode)							Y	N
	INSP1	IA Pressure limit functions	Ask CLI to adjust IA pressure to 40 cm H2O	After 120 seconds, IA pressure at most 40cmH20	±(2 +(4 % of the actual reading)) cmH2O		Y	N				M
	INSP2	IA Pressure adjustable	Ask CLI to adjust IA pressure to 15, 20, 25, 30, 35 cmH2O	After each reset, IA pressure limited got configuration (test pressure 60 seconds)			Y	N				M
	INSP3	IA Failsafe	Cough, blow, or increase pressure by striking bag representing patient	Observe that failsafe valve opens and that pressure does not exceed 80 cm H2O	±(2 +(4 % of the actual reading)) cmH2O		Y	N				M
	INSP4	IA Failsafe, machine cycle	Repeat INSP3 on machine min, nominal and max rate and volume settings	Observe that failsafe valve opens and that pressure does not exceed 80 cm H2O	±(2 +(4 % of the actual reading)) cmH2O		Y	N				M
5							Y	N
6							Y	N
7							Y	N
8							Y	N
9							Y	N
10							Y	N
11							Y	N
12							Y	N
13							Y	N
14							Y	N
15							Y	N
16							Y	N
17							Y	N
18							Y	N
19							Y	N
20							Y	N
PEEP	(PEEP may be implement by separate external valve. Consider separate part to be part of V for purpose of test)						Y	N
	PEEP1	PEEP is understandable	Ask CLI to explain how to adjust PEEP	They correctly explain			Y	N				M
	PEEP2	LOW PEEP	Ask CLI to adjust PEEP to 5cmH2O	Within 120 seconds, pressure in airway never drops below 5cmH20	±(2 +(4 % of the actual reading)) cmH2O		Y	N				M
	PEEP3	Incremental PEEP	Ask CLI To adjust PEEP to 10, 15, and 20 cmH2O	Within 120 seconds, pressure in airway never drops below 5cmH20	±(2 +(4 % of the actual reading)) cmH2O		Y	N				M
4							Y	N
5							Y	N
6							Y	N
7							Y	N
8							Y	N
9							Y	N
10							Y	N
11							Y	N
12							Y	N
13							Y	N
14							Y	N
15							Y	N
16							Y	N
17							Y	N
18							Y	N
19							Y	N
20							Y	N
Respiration Rate							Y	N
	RAT1	Rate	Measure respirations per minute	Matches DRATE, persistent rate, or physical knobs	1 rpm + 10% of setting		Y	N				M
	RAT2	Up to 30 rpm	Ask CLI to configure machine for 30 rpm. When they are done, measure rpms.	Within 60 seconds, CLI can configure machine to 30 rpm, and V so performs.	2 rpm		Y	N				M
	RAT3	Down to 10 rpm	Ask CLI to configure machine for 10 rpm. When they are done, measure rpms.	Within 60 seconds, CLI can configure machine to 10 rpm, and V so performs.	2 rpm		Y	N				M
	RAT4	Increments	Ask CLI to configure machine to 22 rpm. When they are done, measure rpms.	Within 60 seconds, CLI can configure machine to 22 rpm, and V so performs.	2 rpm		Y	N				M
5							Y	N
6							Y	N
7							Y	N
8							Y	N
9							Y	N
10							Y	N
11							Y	N
12							Y	N
13							Y	N
14							Y	N
15							Y	N
16							Y	N
17							Y	N
18							Y	N
19							Y	N
20							Y	N
Tidal Volume							Y	N
	TID1	400ml Vt functions	Ask CLI to configure machine to 400ml Tidal Volume (if only setting for machine, ask them to state.) When so set, measure with professional plastic test lung. Observe rise of lung(measured in mm) to match that of forcing 400ml into lung via syringe or other precise method.)	400ml	±10%		Y	N				M
	TID2	Optional 350 and 450ml setting	Ask CLI to configure machine to 350ml and 450ml Tidal Volume (if only setting for machine, ask them to state.) When so set, measure with professional plastic test lung. Observe rise of lung(measured in mm) to match that of forcing 400ml into lung via syringe or other precise method.)	350ml and 450ml	±10%		Y	N				O
	TID3	Optional 250 and 600ml range	Ask CLI to configure machine to extremes of range 250ml and 600ml Tidal Volume (if only setting for machine, ask them to state.) When so set, measure with professional plastic test lung. Observe rise of lung(measured in mm) to match that of forcing 400ml into lung via syringe or other precise method.)	250ml and 600ml	±10%		Y	N				O
	TID4	Highly Optional 800ml	Ask CLI to configure machine to extremes of range 800ml Tidal Volume (if only setting for machine, ask them to state.) When so set, measure with professional plastic test lung. Observe rise of lung(measured in mm) to match that of forcing 400ml into lung via syringe or other precise method.)	800ml	±10%		Y	N				O
5							Y	N
6							Y	N
7							Y	N
8							Y	N
9							Y	N
10							Y	N
11							Y	N
12							Y	N
13							Y	N
14							Y	N
15							Y	N
16							Y	N
17							Y	N
18							Y	N
19							Y	N
20							Y	N
Infection Control							Y	N
	INFC1	Device can be fully implemented while wearing PPE	Ask CLI to fully connect the system, turn it on, and verify function	CLI can successfuly execute CON1-7 and SUP1-6 while wearing full PPE including hood			Y	N				M
	INFC2	Surfaces can be cleaned	Ask CLI to disinfect external surfaces	Within 60 seconds, CLI can disinfect external parts of device	none		Y	N				M
	INFC3	HMEF-BV can be changed	Ask CLI to change and dispose of HMEF	Within 30 seconds, CLI can remove HMEF and replace it	none		Y	N				M
	INFC4	Room air inlet/outlet filter can be changed	Ask CLI to change room air inlet/outlet filters	Within 3 minutes, CLI can change room air inlet/outlet filters	<30 seconds		Y	N				M
7							Y	N
8							Y	N
9							Y	N
10							Y	N
11							Y	N
12							Y	N
13							Y	N
14							Y	N
15							Y	N
16							Y	N
17							Y	N
18							Y	N
19							Y	N
20							Y	N
Monitoring							Y	N
	MON1	Vt verified	Ask CLI to verify and report Vt	Within 5 seconds, CLI can report set and actual Vt from monitor	<2 seconds		Y	N				M
	MON2	PEEP verified	Ask CLI to verify and report PEEP	Within 5 seconds, CLI can report set and achieved PEEP from monitor	<2 seconds		Y	N				M
	MON3	DFiO2 verified	Ask CLI to verify and report DFiO2	Within 5 seconds, CLI can report set and achieved FiO2 from monitor	<2 seconds		Y	N				M
	MON4	Breathing rate verified	Ask CLI to verify and report breathing rate	Within 5 seconds, CLI can report patient breathing rate	<2 seconds		Y	N				M
	MON5	Airway pressure	Ask CLI to verify and report airway pressure	Within 5 seconds, CLI can report airway pressure	<2 seconds		Y	N				M
	MON6	CO2 verified	Ask CLI to verify and report CO2	Within 5 seconds, CLI can report patient expiratory CO2	<2 seconds		Y	N				O
	MON7	Real-time patient breath verified	Ask CLI to verify and report last patient breath	Within 5 seconds, CLI can report last patient breath in real-time	<2 seconds		Y	N				M
	MON8	I:E ratio verified	Ask CLI to verify and report I:E ratio	Within 5 seconds, CLI can report set I:E ratio	<2 seconds		Y	N				O
	MON9	Battery life remaining verified	Ask CLI to verify battery life remaining	Within 5 seconds, CLI can correctly identify battery life	<2 seconds		Y	N				O
10							Y	N
11							Y	N
12							Y	N
13							Y	N
14							Y	N
15							Y	N
16							Y	N
17							Y	N
18							Y	N
19							Y	N
20							Y	N
Alarm							Y	N
	ALA1	Device power failure	Instruct CLI to cut power to the device without proper shut off protocol	Alarm is immediately set off			Y	N				M
	ALA2	Pressure limit exceeded	Instruct CLI to forcibly add pressure exceeding 80cm H20 whilee blocking failsafe valve	Alarm is set off after 3 seconds	<2 seconds		Y	N				M
	ALA3	Oxygen supply failure	Instruct CLI to disconnect oxygen supply	Alarm is set off after 5 seconds	<5 seconds		Y	N				M
	ALA4	Patient breath below acceptable range	Instruct CLI to block ETT	Alarm is set off after 3 seconds	<2 seconds		Y	N				M
	ALA5	Hypercapnia	Instruct CLI to hold breath for 10 seconds and then hyperventilate into device	Alarm is set off after 3 seconds	<2 seconds		Y	N				O
	ALA6	Hypocapnia	Instruct CLI to hyperventilate for 10 seconds and then breathe into device	Alarm is set off after 3 seconds	<2 seconds		Y	N				O
	ALA7	PEEP failure	Instruct CLI to set PEEP and unplug ETT at end of exhalation cycle	Alarm is set off after 2 seconds	<1 second		Y	N				M
	ALA8	IA pressure failure	Instruct CLI to set IA pressure and unplug ETT during inspiration cycle	Alarm is set off after 2 seconds	<1 second		Y	N				M
	ALA9	Vt not achieved					Y	N				M
	ALA10	Vt exceeded					Y	N				O
	ALA11	Change to battery power	Instruct CLI to cut power to the device, thus switching the device to battery power	Alarm immediately set off			Y	N				O
	ALA12	Low battery	Instruct CLI to cut power to the device, thus switching the device to battery power, and wait at least 15min.	Alarm is set off	<1min		Y	N				O
	ALA13	Battery power failure	Instruct CLI to cut power to the device, thus switching the device to battery power, and wait for battery power to fail.	Alarm is set off			Y	N				M
14							Y	N
15							Y	N
16							Y	N
17							Y	N
18							Y	N
19							Y	N
20							Y	N
Inspiration:Expiration Ratio							Y	N				M
	IER1	Inspiratory:Expiratory 1:2 ratio functions	Ask CLI to configure 1:2 I:E (or state the ratio if only available setting). Once set, connect to test lung and verify that inspiratory:expiratory times match a 1:2 second ratio for the set breathing rate.	Timing ratio matches	<0.1s		Y	N				O
	IER2	Inspiratory:Expiratory ratio can be changed	Ask CLI to change I:E ratio by incrementing and decrementing expiratory time (changes of 0.1) from 1:1 to 1:4, and measuring with a test lung. Measure the expiratory time to ensure that I:E ratio matches.	Timing ratio matches	<0.1s		Y	N
Inverse Inspiration:Expiration Ratio							Y	N
	INVIER1	Inspiratory:Expiratory 2:1 ratio functions	Ask CLI to configure 2:1 I:E (or state the ratio if only available setting). Once set, connect to test lung and verify that inspiratory:expiratory times match a 2:1 second ratio for the set breathing rate.				Y	N				O
	INVIER2	Inverse inspiratory:expiratory ratio can be changed	Ask CLI to change I:E ratio by incrementing expiratory time (changes of 0.1) from 2:1 to 10:1, and measuring with a test lung. Measure the expiratory time to ensure that I:E ratio matches.				Y	N				O
Fraction of Inspired Oxygen							Y	N
	FIO1	FiO2 functions	Ask CLI to configure 50% FiO2, measure with an oxygen sensor, and verify that the values correspond	Matches			Y	N				M
	FIO2	FiO2 can be changed	Ask CLI to lower FiO2 to 30%, and increase (10% increments) up to 100%. Verify each step with oxygen sensor.	Matches			Y	N				M
3							Y	N
4							Y	N
5							Y	N
6							Y	N
7							Y	N
8							Y	N
9							Y	N
10							Y	N
11							Y	N
12							Y	N
13							Y	N
14							Y	N
15							Y	N
16							Y	N
17							Y	N
18							Y	N
19							Y	N
20							Y	N
Durability							Y	N
	DUR1	Dropped and still functioning	Ask CLI to drop the device from bed height ten times when not in use. Execute CON1-7, SUP1-6, and verify function.	All functions remain.			Y	N				M
	DUR2	Dropped while deployed and still functioning	Ask CLI to setup device and from 1.5m ten times. Verify function.	All functions remain.			Y	N				O
	DUR3	Can be transported in ambulance-type setting	Ask CLI to setup device and vigorously shake for 30seconds.	All connections remain.			Y	N				O
	DUR4	Can run continously for extended periods of time	Ask TECH to setup device and have it run for 14 days. Continually verify status.	All functions stable throughout timeframe.			Y	N				M
	DUR5	Connections are robust	Ask CLI to connect device and firmly tug each connection	All connections remain stable			Y	N				O
	DUR6	Water resistant	Tilt V an angle up to 15 degrees from its normal position and exposed to dripping water for a / duration of 10 minutes and a water flow equivalent to 3mm rainfall per minute (IP22)	All functions remain			Y	N				O
7							Y	N
8							Y	N
9							Y	N
10							Y	N
11							Y	N
12							Y	N
13							Y	N
14							Y	N
15							Y	N
16							Y	N
17							Y	N
18							Y	N
19							Y	N
20							Y	N
Labelling							Y	N
	LABL1	Device COVID-19 restriction	Ask CLI to identify label stating use only for COVID-19	CLI can identify and read out loud the label stating "Restricted device for use during COVID-19 pandemic, only to be used for emergency ventilation" within 20 seconds	<10 seconds		Y	N				M
	LABL2	Functions and controls easily identified	Ask CLI to identify all controls and functions on device	CLI can correctly identify all controls and functions of the device in 1min	<20 seconds		Y	N				M
	LABL3	Breathing inlet and outlet easily identified	Ask CLI to identify breathing inlet/outlet	CLI can correctly identify inlet/outlet label within 20 seconds	<10 seconds		Y	N				M
	LABL4	Manual backup label easily identified	Ask CLI to identify label stating "Manual Back Up Ventilation Must Be Available"	CLI can correctly identify and read label out loud in 20 seconds	<10 seconds		Y	N				M
	LABL5	Default settings easily identified	Ask CLI to identify and state the default settings based on labels	CLI can correctly identify default settings within 1min	<10 seconds		Y	N				M
	LABL6	(For transient high-flow device) High flow device label easily identified	Ask CLI to identify the "High-Flow Device" label	CLI can correctly identify label within 20 seconds	<10 seconds		Y	N				O
7							Y	N
8							Y	N
9							Y	N
10							Y	N
11							Y	N
12							Y	N
13							Y	N
14							Y	N
15							Y	N
16							Y	N
17							Y	N
18							Y	N
19							Y	N
20							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N
							Y	N

## References
	Important Reference Documents

Reference #
1	Rapidly Manufactured Ventilator System	https://assets.publishing.service.gov.uk/government/uploads/system/uploads/attachment_data/file/874279/RMVS001_Rapidly_Manufactured_Ventilator_Specification__PDF.pdf

	These are different internatinal standards, each of which may drive a test, and may overlap or even conflit. The colum in the table says which standard comes from where.
Standards	Name	Defining document or repo	Notes
	RMVS	ref(1), British	This standard has optional elements we will denote RMVS(O).

Inverse respiratory ventilation for IER tests	https://www.ncbi.nlm.nih.gov/books/NBK535395/	http://www.anaesthesia.med.usyd.edu.au/resources/lectures/ventilation_clt/ventilation.html

Draft v0.1 - March 25th
By Public Invention, in cooperation with BreathOfLife and Jonathan Orban
Licensed under CC0; consider a template to copy and extend
Copyright Robert L. Read, Juan E Villacres Perez, Jonathan Orban and others 2020
Note: primary source is UK Rapid Manufactered Ventilator System (1)
Primary Point of Contact: Robert L. Read, PhD <read.robert@gmail.com>

## Abbreviations
Abbreviations	Definition
CLI	Clinician (Doctor, nurse, respiratory therapist)
DTIDVOL	Default Tidal Volume
DRATE	Default Respiratory Rate
DFiO2	Default Fraction of Inspired Oxygen
FiO2	Fraction of Inspired Oxygen
V	Ventilator
rpm	respirations per minute
IA	Inspiratory Airway
Vt	Tidal Volume
ETT	Endotracheal Tube
DIERATIO	Default Inspiration / Expiration Ratio
IER	Inspiratory:Expiratory ratio
HMEF-BV	Heat and Moisture Exchange Filter (Bacterial and Viral filter)
PPE	Personal Protective Equipment - Complex protective clothing which includes: Eye goggles (in addition to spectacles if worn), Face shield, Plastic apron, Surgical gown, Two layers of gloves, usually nitrile nonhanded small, medium, large variants, Gloves are donned in layers and sticky taped onto sleeves of gown in between layers
TECH	Test technician

Draft v0.1 - March 25th
By Public Invention, in cooperation with BreathOfLife and Jonathan Orban
Licensed under CC0; consider a template to copy and extend
Copyright Robert L. Read, Juan E Villacres Perez, Jonathan Orban and others 2020
Note: primary source is UK Rapid Manufactered Ventilator System (1)
Primary Point of Contact: Robert L. Read, PhD <read.robert@gmail.com>

## Version and Contact Info
Draft v0.1 - March 25th
By Public Invention, in cooperation with BreathOfLife and Jonathan Orban
Licensed under CC0; consider a template to copy and extend
Copyright Robert L. Read, Juan E Villacres Perez, Jonathan Orban and others 2020
Note: primary source is UK Rapid Manufactered Ventilator System (1)
Primary Point of Contact: Robert L. Read, PhD <read.robert@gmail.com>

## Basic User Tests (obsolete, delete later)
				Open-Source Ventilator Tests (Validation/Clinician Tests)
	Draft v0.1 - March 25th
	By Public Invention, in cooperation with BreathOfLife and Jonathan Orban
	Licensed under CC0; consider a template to copy and extend
	Copyright Robert L. Read, Juan E Villacres Perez, Jonathan Orban and others 2020
	Note: primary source is UK Rapid Manufactered Ventilator System (1)
	Primary Point of Contact: Robert L. Read, PhD <read.robert@gmail.com>

	Note: All tests may be performed with an engineer acting as CLI (clinician) as initial tests. However, usability requires using person unfamiliar with machine as CLI.
	Note: Some tests can be partially performed on mock/stub machine to test usability
	Note: Although informed by ref(1), this test is for BreathOfLife and assumes Volume Control as mode of operation

	Test Cases
	Abbreviations:			CLI = Clinician (Doctor, nurse, respiratory therapist)	DTIDVOL = Default Tidal Volume	DRATE = Default Respiration Rate	DFiO2 = Default Fraction of Inspired Oxygen		DIERATIO = Default Inspiration / Expiration Ratio
Connection Tests				V = Ventilator	rpm = respirations per minute	IA = Inspiratory airway	Vt = Tidal Volume		ETT = Endotracheal tube
	Assumptions for this Suite:			Training has occured
Number	Name	Necessity	Standard	Description	Verification Method	Expected Result	Tolerance	Test Result	Test Status	Tester Id	Test Date	Device ID
1	CON1	Mandatory	RMVS	Electrical connection	Instruct CLI to connect system to ventilator	They can in 15 seconds	< 60 seconds
2	CON2	Mandatory		Oxygen connection	Instruct CLI to connect O2 supply to ventilator	They can in 60 seconds	< 90 seconds
3	CON3	Optional		Replace Batteries	Instruct CLI to replace batteries	They can in 5 minutes	< 10 minutes
4	CON4	Mandatory		ETT connection	Instruct CLI to connect ETT (in-situ) to ventilator	They can in 45 seconds	< 60 seconds
5	CON5	Mandatory		Humidifier connection	Instruct CLI to connect humidifier to ventilator	They can in 60 seconds	90 seconds
6	CON6	Mandatory		Inlet and outlet separated by 10cm	Meaure distance.	At least 10 cm	none
7	CON7	Mandatory		Male inlet and outlet ports provide clearance	Check inlet and outputs clear in 5cm circle from center	At least 5cm circle	none

Startup
	Assumptions for this Suite:			Training has occured
Number	Name			Description	Verification Method	Expected Result	Tolerance	Test Result	Test Status	Tester Id	Test Date	Device ID
1	SUP1			CLI can tell if powered	Instruct CLI to plug in, ask CLI if it is powered. Instruct to plug into unconnected extension cord. Ask CLI if powered.	CLI correctly can tell if powered.
2	SUP2			Mandatory Mode on Power-Up	On initial power-up, system enters mandatory mode and begins working	Mandatory mode operating with respiration rate = DRATE and volume = DVOL and
3	SUP3			Volume Control begins correct	< 90 seconds	CLI answere 400 ml and V produced 400ml	20ml?
4	SUP4			Respiration Rate Standard	On power-up, check respiration rate, measure total respirations per minute	Respiration rate = DRATE, or persistent Respiration Rate, or match control setting if tied to knobs	1 rpm?
5	SUP5			I:E ratio corrects	On power-up estimate inspiration and expiration absolute times in tenths of seconds (w/ stopwatch), comptue I:E ratio	I:E ratio = DIERATIO, or remebered setting, or physical setting	10%
6	SUP6			Mode	On power-up, in mandatory mode, or persistent mode if that is a feature	Yes	none
7	SUP7	Mandatory	RMVS	CLI can locate instruction label	Instruct CLI to locate and inspect instruction label	CLI can correctly locate and read the instruction label in 20 seconds	30 seconds

Inspiratory Airway Pressure (Volume Control Mode)
	Assumptions for this Suite:			Training has occured
Number	Name			Description	Verification Method	Expected Result	Tolerance	Test Result	Test Status	Tester Id	Test Date	Device ID
1	INSP1			IA Pressure limit functions	Ask CLI to adjust IA pressure to 40 cm H2O	After 120 seconds, IA pressure at most 40cmH20	2 cm H2O
2	INSP2			IA Pressure adjustable	Ask CLI to adjust IA pressure to 15, 20, 25, 30, 35 cmH2O	After each reset, IA pressure limited got configuration (test pressure 60 seconds)
3	INSP3			IA Failsafe	Cough, blow, or increase pressure by striking bag representing patient	Observe that failsafe valve opens and that pressure does not exceed 80 cm H2O	4 cm H2O?
3	INSP4			IA Failsafe, machine cycle	Repeat INSP3 on various machine extreme rate and volume settings	Observe that failsafe valve opens and that pressure does not exceed 80 cm H2O	4 cm H2O?

PEEP	(PEEP may be implement by separate external valve. Consider separate part to be part of V for purpose of test)
	Assumptions for this Suite:			Training has occured
Number	Name			Description	Verification Method	Expected Result	Tolerance	Test Result	Test Status	Tester Id	Test Date	Device ID
1	PEEP1			PEEP is understandable	Ask CLI to explain how to adjust PEEP	They correctly explain
2	PEEP2			LOW PEEP	Ask CLI to adjust PEEP to 5cmH2O	Within 120 seconds, pressure in airway never drops below 5cmH20	1 cmH2O?
3	PEEP3			Incremental PEEP	Ask CLI To adjust PEEP to 10, 15, and 20 cmH2O	Within 120 seconds, pressure in airway never drops below 5cmH20	1 cmH2O?

Respiration Rate
	Assumptions for this Suite:			Training has occured
Number	Name			Description	Verification Method	Expected Result	Tolerance	Test Result	Test Status	Tester Id	Test Date	Device ID
1	RAT1			Rate	Measure respirations per minute	Matches DRATE, persistent rate, or physical knobs	1 rpm
2	RAT2			Up to 30 rpm	Ask CLI to configure machine for 30 rpm. When they are done, measure rpms.	Within 60 seconds, CLI can configure machine to 30 rpm, and V so performs.	2 rpm
3	RAT3			Down to 10 rpm	Ask CLI to configure machine for 10 rpm. When they are done, measure rpms.	Within 60 seconds, CLI can configure machine to 10 rpm, and V so performs.	2 rpm
4	RAT4			Increments	Ask CLI to configure machine to 22 rpm. When they are done, measure rpms.	Within 120 seconds, CLI can configure machine to 22 rpm, and V so performs.	2 rpm

Tidal Volume
	Assumptions for this Suite:			Training has occured
Number	Name			Description	Verification Method	Expected Result	Tolerance	Test Result	Test Status	Tester Id	Test Date	Device ID
1	TID1			400ml Vt functions	Ask CLI to configure machine to 400ml Tidal Volume (if only setting for machine, ask them to state.) When so set, measure with professional plastic test lung. Observe rise of lunch (measured in mm) to match that of forcing 400ml into lung via syringe or other precise method.)	400ml	10ml
2	TID2			Optional 350 and 450ml setting	Ask CLI to configure machine to 350ml and 450ml Tidal Volume (if only setting for machine, ask them to state.) When so set, measure with professional plastic test lung. Observe rise of lunch (measured in mm) to match that of forcing 400ml into lung via syringe or other precise method.)	350ml and 450ml	10ml
3	TID3			Optional 250 and 600ml range	Ask CLI to configure machine to extremes of range 250ml and 600ml Tidal Volume (if only setting for machine, ask them to state.) When so set, measure with professional plastic test lung. Observe rise of lunch (measured in mm) to match that of forcing 400ml into lung via syringe or other precise method.)	250ml and 600ml	10ml
4	TID4			Highly Optional 800ml	Ask CLI to configure machine to extremes of range 800ml Tidal Volume (if only setting for machine, ask them to state.) When so set, measure with professional plastic test lung. Observe rise of lunch (measured in mm) to match that of forcing 400ml into lung via syringe or other precise method.)	800ml	10ml

Infection Control
	Assumptions for this Suite:			Training has occured
Number	Name			Description	Verification Method	Expected Result	Tolerance	Test Result	Test Status	Tester Id	Test Date	Device ID
1	INFC1			Device can be fully implemented while wearing PPE	Ask CLI to fully connect the system, turn it on, and verify functio	CLI can successfuly execute CON1-7 and SUP1-6 while wearing full PPE including hood
2	INFC2			Parts can be changed/disposed	Ask CLI to change ETT	Within 30 seconds, CLI can remove ETT, dispose of it in appropriate biohazard, and replace it
3	INFC3			Surfaces can be cleaned	Ask CLI to disinfect external surfaces	Within 60 seconds, CLI can disinfect external parts of device
4	INFC4			Seal from aerosol droplets
5	INFC5

Monitoring
	Assumptions for this Suite:			Training has occured
Number	Name			Description	Verification Method	Expected Result	Tolerance	Test Result	Test Status	Tester Id	Test Date	Device ID
1	MON1			Vt verified	Ask CLI to verify Vt	Within 5 seconds, CLI can report ser and actual Vt from monitor
2	MON2			PEEP verified	Ask CLI to verify PEEP	Within 5 seconds, CLI can report set and achieved PEEP from monitor
3	MON3			DFiO2 verified	Ask CLI to verify DFiO2	Within 5 seconds, CLI can report set and achieved FiO2 from monitor
4	MON4			Breathing rate verified	Ask CLI to verify breathing rate	Within 5 seconds, CLI can report patient breathing rate
5	MON5			Plateau pressure verified	Ask CLI to verify plateau pressure	Within 5 seconds, CLI can report plateau pressure
6	MON6			CO2 verified	Ask CLI to verify CO2	Within 5 seconds, CLI can report patient expiratory CO2

Alarm
	Assumptions for this Suite:			Training has occured
Number	Name			Description	Verification Method	Expected Result	Tolerance	Test Result	Test Status	Tester Id	Test Date	Device ID
1	ALA1			Device power failure	Instruct CLI to cut power to the device without proper shut off protocol	Alarm is immediately set off
2	ALA2			Pressure limit exceeded	Instruct CLI to forcibly add pressure exceeding 40cm H20	Alarm is set off after 3 seconds	2 seconds
3	ALA3			Oxygen supply failure	Instruct CLI to reduce oxygen flow SpO2<90%	Alarm is set off after 5 seconds	5 seconds
4	ALA4			Air flow stops	Instruct CLI to block ETT	Alarm is set off after 3 seconds	2 seconds
5	ALA5			Hypercapnia
6	ALA6			PEEP failure	Instruct CLI to set PEEP and unplug ETT at end of exhalation cycle	Alarm is set off after 2 seconds	1 second
7	ALA7			IA pressure failure	Instruct CLI to set IA pressure and unplug ETT during inspiration cycle	Alarm is set off after 2 seconds	1 second

## Notes
Note: All tests may be performed with an engineer acting as CLI (clinician) as initial tests. However, usability requires using person unfamiliar with machine as CLI.
Note: Some tests can be partially performed on mock/stub machine to test usability
Note: Although informed by ref(1), this test is for BreathOfLife and assumes Volume Control as mode of operation

Draft v0.1 - March 25th
By Public Invention, in cooperation with BreathOfLife and Jonathan Orban
Licensed under CC0; consider a template to copy and extend
Copyright Robert L. Read, Juan E Villacres Perez, Jonathan Orban and others 2020
Note: primary source is UK Rapid Manufactered Ventilator System (1)
Primary Point of Contact: Robert L. Read, PhD <read.robert@gmail.com>